# Mass Spectrometry (MS)- Principle, Working, Instrumentation, Steps, Applications

- Mass Spectrometry (MS) is an analytical chemistry technique that helps identify the amount and type of chemicals present in a sample by measuring the mass-to-charge ratio and abundance of gas-phase ions.

- In this instrumental technique, sample is converted to rapidly moving positive ions by electron bombardment and charged particles are separated according to their masses.

- Mass spectrum is a plot of relative abundance against the ratio of mass/charge (m/e).

- These spectra are used to determine the elemental or isotopic signature of a sample, the masses of particles and of molecules, and to elucidate the chemical structures of molecules and other chemical compounds.

## Principle of Mass Spectrometry (MS)

- In this technique, molecules are bombarded with a beam of energetic electrons.

- The molecules are ionized and broken up into many fragments, some of which are positive ions. Each kind of ion has a particular ratio of mass to charge, i.e. m/e ratio (value).

- For most ions, the charge is one and thus, m/e ratio is simply the molecular mass of the ion.

- The ions pass through magnetic and electric fields to reach detector where they are detected and signals are recorded to give a mass spectra.

## Working of Mass Spectrometry (MS)

- In a typical procedure, a sample, which may be solid, liquid, or gas, is ionized, for example by bombarding it with electrons.

- This may cause some of the sample’s molecules to break into charged fragments. These ions are then separated according to their mass-to-charge ratio, typically by accelerating them and subjecting them to an electric or magnetic field:

- Ions of the same mass-to-charge ratio will undergo the same amount of deflection.

- The ions are detected by a mechanism capable of detecting charged particles, such as an electron multiplier. Results are displayed as spectra of the relative abundance of detected ions as a function of the mass-to-charge ratio.

- The atoms or molecules in the sample can be identified by correlating known masses (e.g. an entire molecule) to the identified masses or through a characteristic fragmentation pattern.

## Instrumentation and Steps of Mass Spectrometry (MS)

A. Sample Inlet

- Sample stored in large reservoir from which molecules reaches ionization chamber at low pressure in steady stream by a pinhole called “Molecular leak”.

B. Ionization

- Atoms are ionized by knocking one or more electrons off to give positive ions by bombardment with a stream of electrons. Most of the positive ions formed will carry charge of +1.

- Ionization can be achieved by :

  - Electron Ionization (EI-MS)

  - Chemical Ionization (CI-MS)

  - Desorption Technique (FAB)

C. Acceleration

- Ions are accelerated so that they all have same kinetic energy.

- Positive ions pass through 3 slits with voltage in decreasing order.

- Middle slit carries intermediate and finals at zero volts.

D. Deflection

- Ions are deflected by a magnetic field due to difference in their masses.

- The lighter the mass, more they are deflected.

- It also depends upon the no. of +ve charge an ion is carrying; the more +ve charge, more it will be deflected.

E. Detection

- The beam of ions passing through the mass analyzer is detected by detector on the basis of m/e ratio.

- When an ion hit the metal box, charge is neutralized by an electron jumping from metal on to the ion.

- Types of analyzers:

  - Magnetic sector mass analysers

  - Double focussing analysers

  - Quadrupole mass analysers

  - Time of Flight analysers (TOF)

  - Ion trap analyser

  - Ion cyclotron analyser

## Applications of Mass Spectrometry (MS)

- Environmental monitoring and analysis (soil, water and air pollutants, water quality, etc.)

- Geochemistry – age determination, soil and rock composition, oil and gas surveying

- Chemical and Petrochemical industry – Quality control

- Identify structures of biomolecules, such as carbohydrates, nucleic acids

- Sequence biopolymers such as proteins and oligosaccharides

- Determination of molecular mass of peptides, proteins, and oligonucleotides.

- Monitoring gases in patients breath during surgery.

- Identification of drugs abuse and metabolites of drugs of abuse in blood, urine, and saliva.

- Analyses of aerosol particles.

- Determination of pesticides residues in food